Contents

1. Dymo Sensor Hardware Overview	1

1.1 What’s in the Pack	1

1.2 Ports and Controls	2

1.3 Hardware Specifications	3

1.4 Dymo Operation, Keys and Indicators	4

2 Dymo Sensor – GlobiLab Communication	5

2.1 USB Communication	5

2.2 Bluetooth Wireless Communication	5

2.2.1 Set the Dymo to “pairing mode”	5

2.2.2 Pairing with Windows OS	6

2.2.3 Pairing with a Mac OS	6

2.2.4 Pairing with an iPad	7

# 1. Dymo Sensor Hardware Overview

## 1.1 What’s in the Pack

① Dymo sensor ② USB cable ③ Bumper head ④ Warranty page ⑤ Quick Start Guide 1; 2; 3. 4; 5.

5

4

1.2 Ports and Controls

1

The picture below reviews the Dymo ports, controls and indicators:

①  Bluetooth key and LED

②  USB port

③  Lab stand locking screw

④  Lab stand hole

⑤  Screw for locking to a trolley

⑥  Charger input

⑦  On/Off key and LED

7

1

3

4

2

8

6

5

⑧  Hook

## 1.3 Hardware Specifications

| Parameter | Specification |
| --- | --- |
| Measurement | Force, 3-axis acceleration |
| Force Range | ± 50 N |
| Acceleration Range | ± 8 g |
| Sampling Rate | 500 Hz |
| Transmission Rate | - PC/MAC: 25 times per second - iPad: 10 times per second |
| Onboard Memory | None |
| GlobiLab Software Communication Interface | - Bluetooth - USB cable |
| External Power Supply | 6V AC/DC adapter (use Labdisc adapter) |
| Internal Battery Type | Single 3.7 V Li-PO cell |
| Push Buttons | - On/Off button: On, Run, Stop (short press), Off (long press) - Bluetooth button (long press for pairing) |
| Communication LED | LED (Blue) |
| Working Mode LED | Bi-color LED (Red, Green) |

## 1.4 Dymo Operation, Keys and Indicators

| Operation | Key | LED indicator |
| --- | --- | --- |
| Turn Dymo on | Press On/Off key when Dymo is Off | On/Off LED flashes green |
| Turn Dymo off | Press On/Off key for 3 seconds when Dymo is On | Both LED off |
| Start logging data | Press the On/Off key when the Dymo is On | On/Off LED constant green light |
| Stop logging date | Press On/Off key when Dymo is logging data | On/Off LED flashes green |
| Bluetooth pairing | Press Bluetooth key for 3 secons | Bluetooth LED flashes 3 times in blue, then turns off. |
| Zero Force Offset | Turn Off the sensor. Remove any weight from the Force sensor. Disconnect it from the computer. Turn the sensor On. Press BOTH keys for 3 seconds. | After both keys are pressed for 3 seconds, Bluetooth LED lights in constant blue |
| Battery charging | N/A | On/Off LED slow flashing red. |
| Low battery | N/A | On/Off LED fast flashing red light |

# 2 Dymo Sensor – GlobiLab Communication

## 2.1 USB Communication

Upon USB cable connection between the computer and the Dymo, the GlobiLab software automatically detects the USB connection and begins communicating with the sensor.

## 2.2 Bluetooth Wireless Communication

Before wireless communication with a Dymo for the first time, the Dymo should be added as a device to the computer in a process called pairing. Pairing need be done only once for the Dymo, after which the computer stores the connection information, including a unique name for each Dymo sensor. When no Dymo is connected via USB, the computer will automatically try to wirelessly connect to the last connected Dymo. To connect to a different or a new Dymo, right click on the Bluetooth icon in the GlobiLab status bar,  located at the bottom right corner of the screen, then click on the Dymo you want to connect to.

### 2.2.1 Set the Dymo to “pairing mode”

- Turn on the Dymo.

- Press the Dymo Pairing button for at least three seconds until the Pairing LED flashes green three times. The Dymo is now in Bluetooth mode.

### 2.2.2 Pairing with Windows OS

- Start the GlobiLab software.

- Right click the Bluetooth   icon, located on the status bar at bottom right corner of the GlobiLab.

- This action opens a pop-up menu. From the menu select “find more Labdisc and sensors”. The computer opens the “add a device” dialog box and starts searching for the Bluetooth device.

- Your Dymo will be displayed as “Dymo_xxxx” where “xxxx” are the 4 last digits of the Dymo serial number sticker. Select this device and press Next.

- Select the 2nd option: “Enter the device pairing code” and press Next.

- In the next dialog box enter “1234” as the pairing code and click Next. Wait for the computer to finish the process and announce that the Dymo and computer were paired successfully.

- Right click the Bluetooth icon on the status bar. Choose the Dymo you’ve just paired and click on it. The Computer will connect to the Dymo and will turn the Bluetooth icon blue.

### 2.2.3 Pairing with a Mac OS

- Open the Bluetooth menu from the Mac menu bar and select "Set Up Bluetooth Device...".

- A dialog box opens. Your Dymo will be displayed as “Dymo_xxxx” where “xxxx” are the 4 last digits of the Dymo serial number sticker. Select this device and press Continue.

- Put your Dymo in pairing mode again (see 2.2.1)

- In the next dialog box press the “Passcode Options…” button and select the 3rd option: “Use a specific passcode”. Enter the passcode “1234”, press “OK” and “Continue”.

- When the wizard action is complete wait for the computer to confirm that the Dymo was added successfully. The Dymo is now paired, and the software can connect to it.

- Open the GlobiLab software.

- Right click the Bluetooth icon  located at the bottom right corner of the software.

- Choose the Dymo you’ve just paired and click on it. The computer will connect to the Dymo and turn the Bluetooth icon blue.

### 2.2.4 Pairing with an iPad

- Launch the iPad Setting

- Open Bluetooth. Make sure the iPad Bluetooth is on.

- From the devices list click the Dymo_xxxx, where the xxxx digits match the last 4 digits of your Dymo serial number sticker on the Dymo back cover.

- The iPad will ask for a PIN code. Enter “1234” and click Pair.

- The iPad will show Dymo_xxxx connected.

Rev: 31.01.17